## Sheet1
advanced | available | default | description | fbnr | group0 | group1 | group2 | highly secured | longname | max | min | name | parameter | poll | process | read | secured | varlength | vartype | vartype2 | write
FALSE | TRUE | 7SN999999 | identnr.+softwareversion[+serialnr.] | 0 | 13 |  |  | FALSE | Identification string |  |  | identstrng | 1 | FALSE | 0 | TRUE | FALSE | -2 | c |  | TRUE
FALSE | TRUE | 0 | primary node address: network parameter FLOW-BUS | 1 | 1 |  |  | TRUE | Primary node address | 128 | 0 | pna | 2 | FALSE | 0 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | secundary node address: network parameter FLOW-BUS | 2 | 1 |  |  | TRUE | Secondary node address | 128 | 0 | sna | 3 | FALSE | 0 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 1 | next node address: network parameter FLOW-BUS | 3 | 1 |  |  | FALSE | Next node address | 128 | 0 | nna | 4 | FALSE | 0 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 32 | last node address: network parameter FLOW-BUS | 4 | 1 |  |  | FALSE | Last node address | 128 | 0 | lna | 5 | FALSE | 0 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 67 | FLOW-BUS arbitrage setting and/or automatic optimization | 5 | 1 |  |  | TRUE | Arbitrage | 255 | 0 | arbitrage | 6 | FALSE | 0 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE |  | init and reset security key commands for network/parameter settings | 10 | 12 |  |  | FALSE | Initreset | 255 | 0 | initreset | 7 | FALSE | 0 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | measured value (100% = 32000) | 0 | 2 |  |  | FALSE | Measure | 41942 | -23593 | measure | 8 | TRUE |  | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | setpoint: wanted value (100% = 32000) | 1 | 2 | 18 |  | FALSE | Setpoint | 32767 | 0 | setpoint | 9 | TRUE |  | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | setpoint ramp signal 0..100 % in up to slope x 0.1 sec. | 2 | 18 |  |  | FALSE | Setpoint slope | 30000 | 0 | setpslope | 10 | FALSE |  | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | analog input signal, normally used for ext. setp. (100% = 32000) | 3 | 2 | 18 |  | FALSE | Analog input | 41942 | -23593 | analoginp | 11 | TRUE |  | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | 0 | control mode selection for instrument or module | 4 | 18 |  |  | FALSE | Control mode | 255 | 0 | cntrlmode | 12 | FALSE |  | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0.0 | polynomial constant A (offset) | 5 | 3 |  |  | FALSE | Polynomial constant A | 3.4028e+38 | -3.4028e+38 | polycnst A | 13 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1.0 | polynomial constant B (span] | 6 | 3 |  |  | FALSE | Polynomial constant B | 3.4028e+38 | -3.4028e+38 | polycnst B | 14 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0.0 | polynomial constant C | 7 | 3 |  |  | FALSE | Polynomial constant C | 3.4028e+38 | -3.4028e+38 | polycnst C | 15 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0.0 | polynomial constant D | 8 | 3 |  |  | FALSE | Polynomial constant D | 3.4028e+38 | -3.4028e+38 | polycnst D | 16 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0.0 | polynomial constant E (offset) for setpoint or power value | 9 | 3 |  |  | FALSE | Polynomial constant E | 3.4028e+38 | -3.4028e+38 | polycnst E | 17 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1.0 | polynomial constant F (span) for setpoint or power value | 10 | 3 |  |  | FALSE | Polynomial constant F | 3.4028e+38 | -3.4028e+38 | polycnst F | 18 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0.0 | polynomial constant G for setpoint or power value | 11 | 3 |  |  | FALSE | Polynomial constant G | 3.4028e+38 | -3.4028e+38 | polycnst G | 19 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0.0 | polynomial constant H for setpoint or power value | 12 | 3 |  |  | FALSE | Polynomial constant H | 3.4028e+38 | -3.4028e+38 | polycnst H | 20 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1.0 | readout value at 100% in capacity (readout) unit | 13 | 3 | 19 |  | FALSE | Capacity 100% | 3.4028e+38 | -3.4028e+38 | capacity | 21 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 3 | sensor type information for actual reading and sensor/controller indication | 14 | 3 |  |  | FALSE | Sensor type | 255 | 0 | sensortype | 22 | FALSE |  | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | pointer to capacity (readout) unit (classic unit table] | 15 | 3 | 19 |  | FALSE | Capacity unit index | 255 | 0 | capunit | 23 | FALSE |  | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | fluid number: pointer to fluidset (for e.g. polynome, name and cap.) | 16 | 3 |  |  | FALSE | Fluidset index | 8 | 0 | fluidnr | 24 | FALSE |  | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | AIR | name of fluid | 17 | 3 |  |  | FALSE | Fluid name |  |  | fluidname | 25 | FALSE |  | TRUE | TRUE | 10 | c |  | TRUE
TRUE | TRUE | 0 | node address of module with operation rights | 18 | 12 |  |  | FALSE | Claim node | 128 | 0 | claimnode | 26 | FALSE |  | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed) | 19 | 12 |  |  | FALSE | Modify | 255 | 0 | modify | 27 | FALSE |  | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | status information of several alarms/errors in the instrument | 20 | 4 |  |  | FALSE | Alarm info | 255 | 0 | alarminfo | 28 | TRUE |  | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 32 | amount of channels which can be operated | 12 | 17 |  |  | FALSE | Channel amount | 120 | 1 | chanamount | 29 | FALSE | 0 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 1 | first channel that can be operated | 13 | 17 |  |  | FALSE | First channel | 120 | 1 | firstchan | 30 | FALSE | 0 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 32 | last channel that can be operated | 14 | 17 |  |  | FALSE | Last channel | 120 | 1 | lastchan | 31 | FALSE | 0 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | operation by HOST computer enable flag | 1 | 5 |  |  | FALSE | <hostcontrl> | 1 | 0 | hostcontrl | 32 | FALSE | 9 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE |  | alarm message string with unit type information | 0 | 5 |  |  | FALSE | Alarm message unit type |  |  | alrmmsgTA | 33 | TRUE | 10 | TRUE | FALSE | 16 | c |  | FALSE
TRUE | TRUE |  | alarm message string with unit number information | 1 | 5 |  |  | FALSE | Alarm message number |  |  | alrmmsgnr | 34 | TRUE | 10 | TRUE | FALSE | 16 | c |  | FALSE
FALSE | TRUE |  | status of relays/potential free contacts ('0' = not activated, '1' = activated) | 2 | 5 |  |  | FALSE | Relay status |  |  | relstatus | 35 | FALSE | 10 | TRUE | FALSE | 8 | c |  | FALSE
FALSE | TRUE | 0 | actual value of counter | 0 | 5 |  |  | FALSE | Actual counter value | 3.4028e+38 | 0 | actualval | 36 | TRUE |  | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE |  | signal input selection (' '=no value,'+'=pos value,'-'=neg value input) | 1 | 5 |  |  | FALSE | Signal input selection |  |  | signinpsel | 37 | FALSE |  | TRUE | FALSE | 8 | c |  | TRUE
FALSE | TRUE |  | external reset input enable/disable ('E'=enable,' '=disable) | 2 | 5 |  |  | FALSE | Reset input selection |  |  | resinpsel | 38 | FALSE |  | TRUE | FALSE | 8 | c |  | TRUE
TRUE | TRUE |  | limit/batch for counter in sensor standard units | 3 | 5 |  |  | FALSE | <limit> | 3.4028e+38 | 0 | limit | 39 | FALSE |  | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 00000000 | delay time string in days,hours,minutes,seconds | 4 | 5 |  |  | FALSE | Delay time | 99235959 | 0 | delaytime | 40 | FALSE |  | TRUE | FALSE | 8 | c |  | TRUE
FALSE | TRUE | 00000000 | duration time string in days,hours,minutes,seconds | 5 | 5 |  |  | FALSE | Duration time | 99235959 | 0 | duratntime | 41 | FALSE |  | TRUE | FALSE | 8 | c |  | TRUE
FALSE | TRUE | 00000000 | valve output setting ('0'=do nothing, '1'=close valve) | 6 | 5 |  |  | FALSE | Valve output setting |  |  | vlvoutset | 42 | FALSE |  | TRUE | FALSE | 8 | c |  | TRUE
FALSE | TRUE | LLLLLLLL | relay output setting ('L'=low, 'H'=high, 'P'=pulse (1 sec.)) | 7 | 5 |  |  | FALSE | Relay output setting |  |  | reloutset | 43 | FALSE |  | TRUE | FALSE | 8 | c |  | TRUE
FALSE | TRUE | 0 | operation mode of T/A module | 8 | 5 |  |  | FALSE | Operation mode T/A | 9 | 0 | opermodeTA | 44 | FALSE |  | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | ln/min | readout unit string | 9 | 5 |  |  | FALSE | Readout unit |  |  | readunit | 45 | FALSE |  | TRUE | FALSE | 7 | c |  | FALSE
FALSE | TRUE | 1 | readout factor matching readout unit string | 10 | 5 |  |  | FALSE | Readout factor | 10000000000 | 1e-10 | readfact | 46 | FALSE |  | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE |  | reset unit command (1=reset T/A unit) | 12 | 5 |  |  | FALSE | Reset unit | 1 | 0 | resetunit | 47 | FALSE |  | FALSE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0.1 | valve output differentiation time constant downwards | 9 | 6 |  |  | FALSE | Valve differentiator down | 3.4028e+38 | 0 | TdValveDn | 48 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0.1 | valve output differentiation time constant upwards | 10 | 6 |  |  | FALSE | Valve differentiator up | 3.4028e+38 | 0 | TdValveUp | 49 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 5.8 | sensor signal differentiation time constant downwards | 11 | 6 |  |  | FALSE | Sensor differentiator down | 3.4028e+38 | 0 | TdSensorDn | 50 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 5.6 | sensor signal differentiation time constant upwards | 12 | 6 |  |  | FALSE | Sensor differentiator up | 3.4028e+38 | 0 | TdSensorUp | 51 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 6 | cycle time * 10 msec. main loop signal processing | 12 | 6 |  |  | FALSE | Cycle time | 255 | 0 | CycleTime | 52 | FALSE | 114 | TRUE | TRUE |  | c | f | TRUE
FALSE | TRUE | 0 | analog mode selection for analog operation | 3 | 10 |  |  | FALSE | Analog mode | 255 | 0 | AnalogMode | 53 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE |  | reference voltage output signal for analog operation | 6 | 10 |  |  | FALSE | Reference voltage | 65535 | 0 | VrefOutput | 54 | FALSE | 116 | TRUE | TRUE |  | i |  | TRUE
FALSE | TRUE | 0 | valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc) | 1 | 8 |  |  | FALSE | Valve output | 16777215 | 0 | ValveOut | 55 | TRUE | 114 | TRUE | FALSE |  | l |  | TRUE
FALSE | TRUE | 0,001 | dynamic display factor for display filter (0=max, 1=min goes with par 57) | 1 | 6 |  |  | FALSE | Dynamic display factor | 1 | 0 | DynDispFct | 56 | FALSE | 117 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0,000001 | static display factor for display filter (0=max, 1=min goes with par 56) | 2 | 6 |  |  | FALSE | Static display factor | 1 | 0 | StaDispFct | 57 | FALSE | 117 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | calibration mode selection (not active until cntrlmode has been set to value 9) | 1 | 7 |  |  | FALSE | Calibration mode | 255 | 0 | CalMode | 58 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 61000 | valve offset: amount of DAC steps within 1 potmeter step | 7 | 8 |  |  | FALSE | Valve offset | 65535 | -32767 | ValveOffst | 59 | FALSE | 116 | TRUE | TRUE |  | i |  | TRUE
FALSE | TRUE | 7 | monitor: output signal (measure) selection for bus and analog output | 2 | 2 |  |  | FALSE | Monitor mode | 255 | 0 | Monitor | 60 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE |  | alarm register containing warning flags | 2 | 4 |  |  | FALSE | Alarm register1 |  |  | AlarmReg1 | 61 | FALSE | 114 | TRUE | TRUE | 8 | c |  | TRUE
TRUE | TRUE |  | alarm register containing critical error flags | 3 | 4 |  |  | FALSE | Alarm register2 |  |  | AlarmReg2 | 62 | FALSE | 114 | TRUE | TRUE | 8 | c |  | TRUE
TRUE | TRUE | 210A7D | calibration register zero scale input 1 ADC | 1 | 9 |  |  | FALSE | <CalRegZS1> | 16777215 | 0 | CalRegZS1 | 63 | FALSE | 116 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 52A513 | calibration register full scale input 1 ADC | 2 | 9 |  |  | FALSE | <CalRegFS1> | 16777215 | 0 | CalRegFS1 | 64 | FALSE | 116 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 210A7D | calibration register zero scale input 2 ADC | 3 | 9 |  |  | FALSE | <CalRegZS2> | 16777215 | 0 | CalRegZS2 | 65 | FALSE | 116 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 52A513 | calibration register full scale input 2 ADC | 4 | 9 |  |  | FALSE | <CalRegFS2> | 16777215 | 0 | CalRegFS2 | 66 | FALSE | 116 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 18904E | ADC control register | 4 | 9 |  |  | FALSE | ADC control register | 4294967295 | 0 | ADCcntrReg | 67 | FALSE | 114 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 0 | sensor bridge zero potmeter setting | 5 | 9 |  |  | FALSE | Bridge potmeter | 255 | 0 | BridgePotm | 68 | FALSE | 116 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 1 | broadcast alarm messsage enable flag | 4 | 4 |  |  | FALSE | <AlarmEnble> | 255 | 0 | AlarmEnble | 69 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | test mode selection (not active until cntrlmode has been set to value 5) | 5 | 4 |  |  | FALSE | Test mode | 255 | 0 | TestMode | 70 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 1 | channel selection ADC | 6 | 9 |  |  | FALSE | <ADC channel select> | 32 | 1 | ChanSelect | 71 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | controller response for normal steps (128=normal, <128=slower, >128=faster) | 5 | 8 |  |  | FALSE | Normal step controller response | 255 | 0 | ContrResp | 72 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 1.0 | analog input filter constant (0=max, 1=min) | 3 | 6 |  |  | FALSE | Setpoint exponential smoothing filter | 1 | 0 | ErrInFilCo | 73 | FALSE | 117 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1.0 | sensor input filter constant (0=max, 1=min) | 4 | 6 |  |  | FALSE | Sensor exponential smoothing filter | 1 | 0 | ExpSmooCon | 74 | FALSE | 117 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 32767 | analog output correction factor zero scale (meas outp DSCM-A 0=0 other 32767=0) | 21 | 10 |  |  | FALSE | Analog output zero scale | 65535 | -32767 | AnOutCorZS | 75 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 2000 | analog output correction factor full scale (meas outp 2000 = 1 * multiplication) | 22 | 10 |  |  | FALSE | Analog output full scale | 65535 | 0 | AnOutCorFS | 76 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 32767 | analog input correction factor zero scale (ext setp DSCM-A 0=0 other 32767=0) | 23 | 10 |  |  | FALSE | Analog input zero scale | 65535 | -32767 | AnInpCorZS | 77 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 2000 | analog input correction factor full scale (ext setp 2000 = 1 * multiplication) | 24 | 10 |  |  | FALSE | Analog input full scale | 65535 | 0 | AnInpCorFS | 78 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | (auto)tuning mode selection (not active until cntrlmode has been set to value 6) | 7 | 7 |  |  | FALSE | Tuning mode | 255 | 0 | TuningMode | 79 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | valve type (needed for controlling behaviour) | 6 | 8 |  |  | FALSE | Valve default | 255 | 0 | DefVlvType | 80 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | contains number(s) of changed processes for indirect polling (0xXX / 0xFF) | 19 | 12 |  |  | FALSE | Global modify | 255 | 0 | GlobModify | 81 | FALSE | 0 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0.1 | correction factor valve curve ratio high/low area | 7 | 8 |  |  | FALSE | Valve span correction factor | 1 | 0 | SpanCorr | 82 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 20,80 | Valve curve correction for controller (max. factor*0.1, flow where factor = 1) | 8 | 8 |  |  | FALSE | Valve curve correction |  |  | VlvCrvStps | 83 | FALSE | 114 | TRUE | TRUE | -2 | c |  | TRUE
TRUE | TRUE | 10,5000,10,5000 | array with memberships for normal Fuzzy controller | 9 | 8 |  |  | FALSE | <MemShipNor> |  |  | MemShipNor | 84 | FALSE | 114 | TRUE | TRUE | -2 | c |  | TRUE
TRUE | TRUE | 1,10000,3750,20000 | array with memberships for 0-open Fuzzy controller | 10 | 8 |  |  | FALSE | <MemShipOpn> |  |  | MemShipOpn | 85 | FALSE | 114 | TRUE | TRUE | -2 | c |  | TRUE
FALSE | TRUE | 4 | IO status byte for jumper settings and LED signal modes | 11 | 12 | 20 |  | FALSE | IO status | 255 | 0 | IOStatus | 86 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | DMFC | (FLOW-BUS) device type information string | 1 | 13 |  |  | FALSE | Device type |  |  | DeviceType | 90 | FALSE | 113 | TRUE | FALSE | 6 | c |  | FALSE
FALSE | TRUE | F201C-FA | BHT model number information string | 2 | 13 |  |  | FALSE | BHT model number |  |  | ModelNum | 91 | FALSE | 113 | TRUE | TRUE | -2 | c |  | TRUE
FALSE | TRUE | SN999999A | serial number information string | 3 | 13 |  |  | TRUE | Serial number |  |  | SerialNum | 92 | FALSE | 113 | TRUE | TRUE | -2 | c |  | TRUE
FALSE | TRUE | STANDARD | customer model number information string | 4 | 11 |  |  | FALSE | Customer model number |  |  | MfrConfig | 93 | FALSE | 113 | TRUE | TRUE | -2 | c |  | TRUE
TRUE | TRUE | 01,01,95 | special BHT parameter (to be changed by BHT only) | 1 | 14 |  |  | TRUE | BHT1 |  |  | BHT1 | 94 | FALSE | 118 | TRUE | TRUE | -2 | c |  | TRUE
TRUE | TRUE | 0 | special BHT parameter | 2 | 14 |  |  | FALSE | BHT2 | 65535 | 0 | BHT2 | 95 | FALSE | 118 | TRUE | FALSE |  | i |  | FALSE
TRUE | TRUE | 0 | special BHT parameter | 3 | 14 |  |  | FALSE | BHT3 | 3000000000 | -3000000000 | BHT3 | 96 | FALSE | 118 | TRUE | FALSE |  | l |  | FALSE
TRUE | TRUE | 0 | special BHT parameter | 4 | 14 |  |  | FALSE | BHT4 | 65535 | 0 | BHT4 | 97 | FALSE | 118 | TRUE | FALSE |  | i |  | FALSE
TRUE | TRUE | 0 | special BHT parameter | 5 | 14 |  |  | FALSE | BHT5 | 255 | 0 | BHT5 | 98 | FALSE | 118 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 0 | special BHT parameter | 6 | 14 |  |  | FALSE | BHT6 | 255 | 0 | BHT6 | 99 | FALSE | 118 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | special BHT parameter | 7 | 14 |  |  | FALSE | BHT7 | 255 | 0 | BHT7 | 100 | FALSE | 118 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 0 | special BHT parameter | 8 | 14 |  |  | FALSE | BHT8 | 255 | 0 | BHT8 | 101 | FALSE | 118 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 0 | special BHT parameter | 9 | 14 |  |  | FALSE | BHT9 | 3000000000 | -3000000000 | BHT9 | 102 | FALSE | 118 | TRUE | FALSE |  | l |  | FALSE
TRUE | TRUE | 0 | special BHT parameter (to be changed by BHT only) | 10 | 14 |  |  | TRUE | BHT10 | 1 | 0 | BHT10 | 103 | FALSE | 118 | FALSE | TRUE |  | c |  | TRUE
TRUE | TRUE | 128 | broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve) | 16 | 8 |  |  | FALSE | Broadcast repeating time | 255 | 0 | PulseHight | 104 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | VX.XX | revision number of firmware | 5 | 13 | 20 |  | FALSE | Firmware version |  |  | Version | 105 | FALSE | 113 | TRUE | FALSE | 6 | c |  | FALSE
TRUE | TRUE | 0 | type of pressure sensor | 9 | 20 |  |  | FALSE | Pressure sensor type | 255 | 0 | PressSensr | 106 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 1013.25 | mbar atmospheric (central) barometer pressure | 8 | 20 |  |  | FALSE | Barometer pressure | 1200 | 0 | BaroPress | 107 | FALSE | 116 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 32767 | analog sensor signal input corr. factor zero scale (DSCM-A 0=0 other 32767=0) | 25 | 10 |  |  | FALSE | Sensor input zero scale | 65535 | -32767 | AnIn1CorZS | 108 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 2000 | analog sensor signal input correction factor full scale (2000=1*multiplication) | 26 | 10 |  |  | FALSE | Sensor input full scale | 65535 | 0 | AnIn1CorFS | 109 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 32767 | analog Vref input correction factor zero scale (DSCM-A 0=0 other 32767=0) | 27 | 10 |  |  | FALSE | Reference voltage input zero scale | 65535 | -32767 | AnIn2CorZS | 110 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 2000 | analog Vref input correction factor full scale (2000=1*multiplication) | 28 | 10 |  |  | FALSE | Reference voltage input full scale | 65535 | 0 | AnIn2CorFS | 111 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 32767 | analog setpoint output correction factor zero scale (DSCM-A 0=0 other 32767=0) | 29 | 10 |  |  | FALSE | Analog setpoint zero scale | 65535 | -32767 | AnOu1CorZS | 112 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 2000 | analog setpoint output correction factor full scale (2000=1*multiplication) | 30 | 10 |  |  | FALSE | Analog setpoint full scale | 65535 | 0 | AnOu1CorFS | 113 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
FALSE | TRUE | 0 | reset facilities (program/alarm/batchcounter) | 8 | 12 |  |  | FALSE | Reset | 255 | 0 | Reset | 114 | FALSE | 115 | FALSE | FALSE |  | c |  | TRUE
FALSE | TRUE | USERTAG | user definable alias string | 6 | 11 |  |  | FALSE | User tag |  |  | UserTag | 115 | FALSE | 113 | TRUE | FALSE | -2 | c |  | TRUE
FALSE | TRUE | 0 | maximum limit for sensor signal to trigger alarm situation | 1 | 15 |  |  | FALSE | Alarm limit maximum | 41600 | 0 | AlrmMaxLim | 116 | FALSE | 97 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | minimum limit for sensor signal to trigger alarm situation | 2 | 15 |  |  | FALSE | Alarm limit minimum | 41600 | 0 | AlrmMinLim | 117 | FALSE | 97 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | alarm mode | 3 | 15 |  |  | FALSE | Alarm mode | 255 | 0 | AlrmMode | 118 | FALSE | 97 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | alarm relais activity mode during alarm situation | 4 | 15 |  |  | FALSE | Alarm output mode | 255 | 0 | AlrmOutMod | 119 | FALSE | 97 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | setpoint change enable during alarm situation | 5 | 15 |  |  | FALSE | Alarm setpoint mode | 1 | 0 | AlrmStpMod | 120 | FALSE | 97 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | new/safe setpoint during alarm situation (until reset) | 6 | 15 |  |  | FALSE | Alarm new setpoint | 32767 | 0 | AlrmNwSetp | 121 | FALSE | 97 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | actual counter value | 1 | 16 |  |  | FALSE | Counter value | 3.4028e+38 | 0 | CntrValue | 122 | TRUE | 104 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | pointer to counter unit (classic counter unit table) | 2 | 16 |  |  | FALSE | Counter unit index | 31 | 0 | CntrUnit | 123 | FALSE | 104 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | counter limit/batch | 3 | 16 |  |  | FALSE | Counter limit | 3.4028e+38 | 0 | CntrLimit | 124 | FALSE | 104 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | counter relais activity mode when limit/batch has been reached | 4 | 16 |  |  | FALSE | Counter output mode | 255 | 0 | CntrOutMod | 125 | FALSE | 104 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | setpoint change enable during counter limit/batch situation (until reset) | 5 | 16 |  |  | FALSE | Counter setpoint mode | 1 | 0 | CntrStpMod | 126 | FALSE | 104 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%) | 6 | 16 |  |  | FALSE | Counter new setpoint | 32767 | 0 | CntrNwSetp | 127 | FALSE | 104 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | ln | counter readout unit (informative only for older devices) | 7 | 16 |  |  | FALSE | Counter unit |  |  | CntrUntstr | 128 | FALSE | 104 | TRUE | FALSE | 4 | c |  | TRUE
FALSE | TRUE | ln/min | capacity readout unit (informative only for older devices) | 31 | 3 | 19 |  | FALSE | Capacity unit |  |  | capunitstr | 129 | FALSE |  | TRUE | TRUE | 7 | c |  | TRUE
FALSE | TRUE | 0 | counter mode | 8 | 16 |  |  | FALSE | Counter mode | 255 | 0 | CntrMode | 130 | FALSE | 104 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | VX.XX | minimum required hardware revision level for firmware version | 7 | 13 |  |  | FALSE | Minimum hardware revision |  |  | HwRev | 131 | FALSE | 113 | TRUE | FALSE | 1 | c |  | FALSE
TRUE | TRUE | 1.0 | readout factor for direct reading (changes with readunit: local on module, R.O.) | 1 | 17 |  |  | FALSE | <RCreadfact> | 10000000000 | 1e-10 | RCreadfact | 132 | FALSE |  | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 1 | channel number for operation | 2 | 17 |  |  | FALSE | <channumber> | 120 | 1 | channumber | 133 | FALSE |  | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | master channel for master-slave operation | 3 | 17 |  |  | FALSE | <masterchan> | 120 | 0 | masterchan | 134 | FALSE |  | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 32000 | RC slave factor | 4 | 17 |  |  | FALSE | <RCslavefct> | 32000 | 0 | RCslavefct | 135 | FALSE |  | TRUE | FALSE |  | i |  | TRUE
TRUE | TRUE | 3 | physical node address for channel number | 5 | 17 |  |  | FALSE | <inputnode> | 128 | 0 | inputnode | 136 | FALSE |  | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 1 | physical process for channel number | 6 | 17 |  |  | FALSE | <inputproc> | 128 | 0 | inputproc | 137 | FALSE |  | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | ln/min | readout unit for direct reading (local variable on module: read only) | 7 | 17 |  |  | FALSE | <RCreadunit> |  |  | RCreadunit | 138 | FALSE |  | TRUE | FALSE | 7 | c |  | FALSE
FALSE | TRUE | 100.0 | slave factor for master slave control (setp = master output * slave factor) | 1 | 18 |  |  | FALSE | Slave factor | 500 | 0 | SlaveFact% | 139 | FALSE | 33 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | reference voltage input for setpoint signal | 2 | 18 |  |  | FALSE | Reference voltage input | 65535 | 0 | VrefInput | 140 | TRUE | 33 | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | 0 | controller response when controller is stable: ¦measure-setpoint¦ < 2% | 17 | 8 |  |  | FALSE | Stable situation controller response | 255 | 0 | RespStable | 141 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 20 | absolute temperature in degrees Celsius | 7 | 19 | 13 |  | FALSE | Temperature | 500 | -250 | temperatur | 142 | TRUE | 33 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 1013.25 | absolute pressure in mbar (FPP, FRM, FTM) or bar | 8 | 19 | 13 |  | FALSE | Pressure | 3.4028e+38 | -3.4028e+38 | pressure | 143 | TRUE | 33 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | time in milliseconds | 9 | 19 |  |  | FALSE | Time | 3.4028e+38 | 0 | time | 144 | TRUE | 33 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 50 | calibrated volume in litres | 10 | 19 |  |  | FALSE | Calibrated volume | 3.4028e+38 | 0 | calvolume | 145 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | pointer to sensor number in calibration tube (FPP) | 16 | 19 |  |  | FALSE | Sensor number | 4 | 0 | sensornr | 146 | TRUE |  | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Piston Prover operation mode (write) and status information (read back) | 10 | 20 | 19 |  | FALSE | Range select | 99 | 0 | rangeselct | 147 | TRUE | 115 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | maximum admitted duration time for specific procedure (in 100 ms) | 2 | 20 |  |  | FALSE | Time out | 30000 | 0 | TimeOut | 148 | FALSE |  | TRUE | TRUE |  | i |  | TRUE
FALSE | TRUE | 0 | frequency in Hz | 9 | 21 |  |  | FALSE | Frequency | 100000 | 0 | frequency | 149 | TRUE | 33 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 42773.4 | For FRM and FTM imp/m3 and for FCM imp/kg | 10 | 20 |  |  | FALSE | Impulses/m3 | 3.4028e+38 | 0 | imp/m3 | 150 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | volume flow referenced to normal conditions i.e. 0 �C, 1013.25 hPa(a) in ln/min | 5 | 19 |  |  | FALSE | Normal volume flow | 3.4028e+38 | 0 | RefVolFlow | 151 | FALSE | 33 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | volume flow at actual conditions in l/min | 6 | 19 |  |  | FALSE | Volume flow | 3.4028e+38 | -3.4028e+38 | volumeflow | 152 | FALSE | 33 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | relative pressure between atmosphere and sensor position | 11 | 19 |  |  | FALSE | Delta-p | 100000 | -100000 | delta-p | 153 | TRUE | 33 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 1 | scaling factor (multiplication) for readout on display (for optimal resolution) | 13 | 21 |  |  | FALSE | <scalefact> | 10000 | 1 | scalefact | 154 | FALSE | 33 | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | SENSOR0 | label with information about stop sensor | 17 | 19 |  |  | FALSE | Sensor name |  |  | sensorname | 155 | FALSE |  | TRUE | TRUE | 10 | c |  | TRUE
TRUE | TRUE | 15 | enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic | 9 | 15 |  |  | FALSE | Reset alarm enable | 255 | 0 | RstAlarmEn | 156 | FALSE | 97 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 7 | enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic | 9 | 16 |  |  | FALSE | Reset counter enable | 255 | 0 | RstCountEn | 157 | FALSE | 104 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 3 | node number of master instrument output signal for a slave | 14 | 18 |  |  | FALSE | Master node | 128 | 1 | MasterNode | 158 | FALSE | 33 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 1 | process number of master instrument output signal for a slave | 15 | 18 |  |  | FALSE | Master process | 128 | 1 | MasterProc | 159 | FALSE | 33 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 3 | node number of instrument to be operated by another module (keyboard/display) | 16 | 18 |  |  | FALSE | Remote instrument node | 128 | 1 | InstrNode | 160 | FALSE | 33 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 1 | process number of instrument to be operated by another module (keyboard/display) | 17 | 18 |  |  | FALSE | Remote instrument process | 128 | 1 | InstrProc | 161 | FALSE | 33 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0.0 | Mnimum value at 0% for special user readout unit | 18 | 3 |  |  | FALSE | Minimum custom range | 3.4028e+38 | -3.4028e+38 | RangeMin | 162 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 100.0 | Maximum value at 100% for special user readout unit | 20 | 3 |  |  | FALSE | Maximum custom range | 3.4028e+38 | -3.4028e+38 | RangeMax | 163 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Relay/TTL output setting (disabled when used by alarm or counter) | 11 | 12 |  |  | FALSE | Relay/TTL output | 255 | 0 | Relay/TTL | 164 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | Controller response when valve opens from zero | 18 | 8 |  |  | FALSE | Open from zero controller response | 255 | 0 | RespOpen0% | 165 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 1 | Controller settings for special purpose | 20 | 8 |  |  | FALSE | Controller features | 255 | 0 | ContrType | 166 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 10 | PID factor Kp | 21 | 8 | 3 |  | FALSE | PID-Kp | 3.4028e+38 | 0 | PIDKp | 167 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0.05 | PID factor Ti | 22 | 8 |  |  | FALSE | PID-Ti | 3.4028e+38 | 0 | PIDTi | 168 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | PID factor Td | 23 | 8 |  |  | FALSE | PID-Td | 3.4028e+38 | 0 | PIDTd | 169 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1.293 | Density of selected fluid in kg/m3 | 21 | 3 |  |  | FALSE | Density | 3.4028e+38 | 0 | Density | 170 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE |  | Number of calibration certificate (last basic calibration) | 8 | 13 | 3 |  | TRUE | Calibration certificate |  |  | CalCertNr | 171 | FALSE | 113 | TRUE | TRUE | -2 | c |  | TRUE
FALSE | TRUE | 19991231 | Date of last (basic) calibration | 9 | 13 | 3 |  | TRUE | Calibration date |  |  | CalDate | 172 | FALSE | 113 | TRUE | TRUE | 8 | c |  | TRUE
FALSE | TRUE | 000000 | Service number for repair/rebuilding/recalibration | 10 | 13 |  |  | TRUE | Service number |  |  | ServiceNr | 173 | FALSE | 113 | TRUE | TRUE | 15 | c |  | TRUE
FALSE | TRUE | 19991231 | Date of last service action | 11 | 13 |  |  | TRUE | Service date |  |  | ServDate | 174 | FALSE | 113 | TRUE | TRUE | 8 | c |  | TRUE
FALSE | TRUE | 7 | Identification number (type) of instrument/device | 12 | 13 |  |  | TRUE | Identification number | 255 | 0 | IdentNr | 175 | FALSE | 113 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | special BHT parameter (to be changed by BHT only) | 11 | 14 |  |  | TRUE | BHT11 | 255 | 0 | BHT11 | 176 | FALSE | 118 | FALSE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | power supply indication in Vdc | 12 | 12 |  |  | FALSE | Power mode | 50 | 0 | PowerMode | 177 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 3 | pressure inlet (upstream) of fluid in bara (for first fluidnr only) | 13 | 13 | 3 |  | FALSE | Pressure inlet | 100000 | -100000 | Pupstream | 178 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | pressure outlet (downstream) of fluid in bara (for first fluidnr only) | 14 | 13 | 3 |  | FALSE | Pressure outlet | 100000 | -100000 | Pdownstrm | 179 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | orifice diameter in mm | 15 | 13 |  |  | FALSE | Orifice | 1000 | 0 | Orifice | 180 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 20 | temperature of fluid through instrument (for first fluidnr only) | 16 | 13 | 3 |  | FALSE | Fluid temperature | 3.4028e+38 | -273.15 | FluidTemp | 181 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 3 | time alarm and reset action will be delayed when alarm limit has been exceeded | 7 | 15 |  |  | FALSE | Alarm delay | 255 | 0 | AlrmDelay | 182 | FALSE | 97 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | capacity of instrument at zero 0% in sensor base units (mostly equal to zero) | 22 | 3 |  |  | FALSE | Capacity 0% | 3.4028e+38 | -3.4028e+38 | capacity0% | 183 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | number of instrument channels available for this device | 18 | 12 |  |  | FALSE | Number of channels | 120 | 0 | NumOfChan | 184 | FALSE | 0 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 5 | function of device | 20 | 12 |  |  | FALSE | Device function | 255 | 0 | DeviceFunc | 185 | FALSE | 0 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 1 | Channel number to scan with real time information (to be set once) | 1 | 4 |  |  | FALSE | Scan channel | 255 | 1 | ScanChan | 186 | FALSE | 123 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 8 | Parameter number to scan with real time information (to be set once) | 3 | 4 |  |  | FALSE | Scan parameter | 255 | 0 | ScanPar | 187 | FALSE | 123 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 50 | Scan interval time in msec between two samples (to be set once) | 4 | 4 |  |  | FALSE | Scan time | 65535 | 0 | ScanTime | 188 | FALSE | 123 | TRUE | FALSE |  | i |  | TRUE
TRUE | TRUE |  | Scanned data with time label (can be readout event by event) | 10 | 22 | 4 |  | FALSE | Scan data |  |  | ScanData | 189 | FALSE | 123 | TRUE | FALSE | -2 | c |  | FALSE
FALSE | TRUE | 0.04 | First-step offset current/voltage for valve when opening from 0% | 24 | 8 |  |  | FALSE | Valve open | 24 | 0 | ValveOpen | 190 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1 | Amount of runs of a piston prover (0 = stability check) | 13 | 20 |  |  | FALSE | Number of runs | 255 | 0 | NrOfRuns | 191 | FALSE | 115 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 10 | Minimum process time of a piston prover in 0.1 seconds | 14 | 20 |  |  | FALSE | Minimum process time | 255 | 0 | MinProTime | 192 | FALSE | 115 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0.0001 | Leak rate piston prover/piston tare value multiplier | 9 | 20 |  |  | FALSE | Leak rate | 3 | 0 | LeakRate | 193 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE |  | Sets instr. in info mode for 1 read-cycle to check available parameter options | 15 | 12 |  |  | FALSE | Mode info request |  |  | ModeInfReq | 194 | FALSE | 115 | TRUE | TRUE | 4 | c |  | TRUE
TRUE | TRUE |  | Gives info about possible values of a mode in an array as result of ModeInfo req | 16 | 12 |  |  | FALSE | Mode info option list |  |  | ModeInfOpt | 195 | FALSE | 115 | TRUE | FALSE | 255 | c |  | FALSE
TRUE | TRUE |  | Gives description about one of the mode options | 17 | 12 |  |  | FALSE | Mode info option description |  |  | ModeInfDes | 196 | FALSE | 115 | TRUE | FALSE | 255 | c |  | FALSE
TRUE | TRUE | 0 | Enables/disables options for calibration device (8 bits for 8 options) | 18 | 20 |  |  | FALSE | Calibrations options | 255 | 0 | CalType | 197 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | Real mass flow in kg/min | 4 | 20 |  |  | FALSE | Mass flow | 3.4028e+38 | -3.4028e+38 | MassFlow | 198 | TRUE | 33 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 2 | Fieldbus1 address (main interface) | 10 | 23 |  |  | FALSE | Fieldbus1 address | 255 | 0 | BusAddress | 199 | FALSE | 125 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 1 | Configuration setting for interface to other bus-systems | 3 | 23 |  |  | TRUE | Interface configuration | 3 | 0 | InterfConf | 200 | FALSE | 125 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 12000000 | Fieldbus1 baudrate (main interface) | 9 | 23 |  |  | FALSE | Fieldbus1 baudrate | 10000000000 | 0 | Baudrate | 201 | FALSE | 125 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE |  | Fieldbus1 diagnose string (main interface) | 20 | 23 |  |  | FALSE | Fieldbus1 diagnostic string |  |  | BusDiagnos | 202 | FALSE | 125 | TRUE | FALSE | -2 | c |  | FALSE
TRUE | TRUE | 10 | Number of vanes for use in a rotor meter | 22 | 20 |  |  | FALSE | Number of vanes | 255 | 0 | NrOfVanes | 203 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | FLOW-BUS | Fieldbus1 name (main interface) | 21 | 23 |  |  | FALSE | Fieldbus1 name |  |  | Fieldbus | 204 | FALSE | 125 | TRUE | FALSE | -2 | c |  | FALSE
FALSE | TRUE | 0 | measured value for direct reading (in capunits, max.= capacity) | 0 | 2 |  |  | FALSE | fMeasure | 3.4028e+38 | -3.4028e+38 | fMeasure | 205 | TRUE | 33 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | setpoint: wanted value for direct reading (in capunits, max.= capacity) | 3 | 2 | 18 |  | FALSE | fSetpoint | 3.4028e+38 | -3.4028e+38 | fSetpoint | 206 | FALSE | 33 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Mass in g | 23 | 20 |  |  | FALSE | Mass | 3.4028e+38 | -3.4028e+38 | Mass | 207 | TRUE | 33 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE |  | Manufacturer Status register (64 diagnostic bits) | 1 | 4 |  |  | FALSE | Manufacturer status register |  |  | Mstatus | 208 | FALSE | 119 | TRUE | FALSE | 8 | c |  | FALSE
TRUE | TRUE |  | Manufacturer Warning register (64 diagnostic bits) | 2 | 4 |  |  | FALSE | Manufacturer warning register |  |  | Mwarning | 209 | FALSE | 119 | TRUE | FALSE | 8 | c |  | FALSE
TRUE | TRUE |  | Manufacturer Error register (64 diagnostic bits) | 3 | 4 |  |  | FALSE | Manufacturer error register |  |  | Merror | 210 | FALSE | 119 | TRUE | FALSE | 8 | c |  | FALSE
TRUE | TRUE |  | Diagnostic history string (contains history of diag codes) | 4 | 4 |  |  | FALSE | Diagnostic history string |  |  | DiagHist | 211 | FALSE | 119 | TRUE | FALSE | -2 | c |  | TRUE
TRUE | TRUE | 0 | Diagnostic mode (0 = diagnostics off, 1 = diagnostics on) | 5 | 4 |  |  | FALSE | Diagnostic mode | 255 | 0 | DiagMode | 212 | FALSE | 119 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all) | 6 | 4 |  |  | FALSE | Manufacturer status enable | 255 | 0 | MStatEnabl | 213 | FALSE | 119 | FALSE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | Analog measure output, zero adjust | 21 | 10 |  |  | FALSE | Analog output zero adjust | 3.4028e+38 | -3.4028e+38 | AnOutZA | 214 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | Analog measure output, span adjust | 22 | 10 |  |  | FALSE | Analog output span adjust | 3.4028e+38 | -3.4028e+38 | AnOutSA | 215 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Analog setpoint input, zero adjust | 23 | 10 |  |  | FALSE | Analog input zero adjust | 3.4028e+38 | -3.4028e+38 | AnInZA | 216 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | Analog setpoint input, span adjust | 24 | 10 |  |  | FALSE | Analog input span adjust | 3.4028e+38 | -3.4028e+38 | AnInSA | 217 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Sensor input, zero adjust | 25 | 10 |  |  | FALSE | Sensor input zero adjust | 3.4028e+38 | -3.4028e+38 | SensInZA | 218 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | Sensor input, span adjust | 26 | 10 |  |  | FALSE | Sensor input span adjust | 3.4028e+38 | -3.4028e+38 | SensInSA | 219 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Sensor temperature input, zero adjust | 27 | 10 |  |  | FALSE | Temperature input zero adjust | 3.4028e+38 | -3.4028e+38 | TempInZA | 220 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | Sensor temperature input, span adjust | 28 | 10 |  |  | FALSE | Temperature input span adjust | 3.4028e+38 | -3.4028e+38 | TempInSA | 221 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1.0 | Sensor input filter adapt setting | 5 | 6 |  |  | FALSE | Adaptive smoothing factor | 1 | 0 | ExpSmooAd | 222 | FALSE | 117 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 32000 | Slope setpoint step. Setpoint step for the given slopetime | 24 | 18 |  |  | FALSE | Slope setpoint step | 32000 | 0 | SlopeSetp | 223 | FALSE | 33 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 1 | Number of samples for average filter | 6 | 6 |  |  | FALSE | Filter length | 255 | 0 | FilterLen | 224 | FALSE | 117 | TRUE | TRUE |  | i |  | TRUE
FALSE | TRUE | 0 | Actual accuracy in current unit | 25 | 2 |  |  | FALSE | Absolute accuracy | 3.4028e+38 | -3.4028e+38 | fAccuracy | 225 | FALSE | 33 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Lookup table for linearisation index (x and y direction) | 26 | 3 |  |  | FALSE | Lookup table index | 255 | 0 | LookI | 226 | FALSE | 33 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Lookup table for linearisation x | 27 | 3 |  |  | FALSE | Lookup table X | 3.4028e+38 | -3.4028e+38 | LookX | 227 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | Lookup table for linearisation y | 28 | 3 |  |  | FALSE | Lookup table Y | 3.4028e+38 | -3.4028e+38 | LookY | 228 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Lookup table for linearisation at certain temperature index (z direction) | 29 | 3 |  |  | FALSE | Lookup table temperature index | 255 | 0 | LookTempI | 229 | FALSE | 33 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | Lookup table for linearisation at certain temperature (z) | 30 | 3 |  |  | FALSE | Lookup table temperature | 3.4028e+38 | -273.15 | LookTemp | 230 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0.2 | Maximum current/voltage for valve | 25 | 8 |  |  | FALSE | Valve maximum | 24 | 0 | ValveMax | 231 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | Valve output mode selection | 26 | 8 |  |  | FALSE | Valve mode | 255 | 0 | ValveMode | 232 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0.96 | Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96) | 27 | 8 |  |  | FALSE | Valve open correction | 1.5 | 0 | VlvOpenCor | 233 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8) | 28 | 8 |  |  | FALSE | Valve zero hold | 1 | 0 | VlvZeroHld | 234 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0.009 | Valve slope time (seconds) | 29 | 8 |  |  | FALSE | Valve slope | 50 | 0 | ValveSlope | 235 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | IFI data dump protocol communication string | 21 | 23 |  |  | FALSE | IFI data |  |  | IFIData | 236 | FALSE | 0 | TRUE | FALSE | -2 | c |  | TRUE
FALSE | TRUE | 0 | Piston Prover information about used sensors | 20 | 20 | 19 |  | FALSE | Range used | 99 | 0 | RangeUsed | 237 | TRUE | 115 | TRUE | FALSE | -2 | c |  | FALSE
TRUE | TRUE | 0 | Fluidset properties | 31 | 3 |  |  | TRUE | Fluidset properties | 255 | 0 | FldSetProp | 238 | FALSE | 33 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Lookup table unit type | 12 | 3 |  |  | TRUE | Lookup table unit type index | 255 | 0 | LUnitType | 239 | FALSE | 33 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE |  | Lookup table unit type name | 13 | 3 |  |  | TRUE | Lookup table unit type |  |  | LUnTypName | 240 | FALSE | 33 | TRUE | TRUE | 20 | c |  | TRUE
TRUE | TRUE | 0 | Lookup table unit (unit LUTy) | 16 | 3 |  |  | TRUE | Lookup table unit index | 255 | 0 | LUnit | 241 | FALSE | 33 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | kg/s | Lookup table unit name | 17 | 3 |  |  | TRUE | Lookup table unit |  |  | LUnitName | 242 | FALSE | 33 | TRUE | TRUE | 7 | c |  | TRUE
FALSE | TRUE | 0 | Capacity (readout) unit type | 29 | 3 |  |  | FALSE | Capacity unit type index | 255 | 0 | CUnitType | 243 | FALSE |  | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE |  | Capacity (readout) unit type name | 30 | 3 |  |  | FALSE | Capacity unit type |  |  | CUnTypName | 244 | FALSE |  | TRUE | TRUE | 20 | c |  | TRUE
FALSE | TRUE | 0 | Capacity (readout) unit type temperature (�C) | 10 | 3 |  |  | FALSE | Capacity unit type temperature | 3.4028e+38 | -273.15 | CUnTypTemp | 245 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1 | Capacity (readout) unit type pressure (bar (a)) | 11 | 3 |  |  | FALSE | Capacity unit type pressure | 3.4028e+38 | 0 | CUnTypPres | 246 | FALSE | 33 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | Minimum capacity in output capacity units | 27 | 3 |  |  | FALSE | Capacity minimum | 3.4028e+38 | -3.4028e+38 | CapMin | 247 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1 | Maximum capacity in output capacity units | 28 | 3 |  |  | FALSE | Capacity maximum | 3.4028e+38 | -3.4028e+38 | CapMax | 248 | FALSE |  | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Formula type needed for conversion | 17 | 3 |  |  | FALSE | Formula type | 65535 | 0 | FormulaTyp | 249 | FALSE | 113 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Heat capacity (Cp) (sensor conditions) | 18 | 3 |  |  | FALSE | Heat capacity | 3.4028e+38 | 0 | HeatCap | 250 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Thermal conductivity (sensor conditions) | 20 | 3 |  |  | FALSE | Thermal conductivity | 3.4028e+38 | 0 | ThermCond | 251 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Dynamic viscosity (fluid conditions) | 21 | 3 |  |  | FALSE | Viscosity | 3.4028e+38 | 0 | Viscosity | 252 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | Standard mass flow in ln/min (20�C, 1.01325 bar (a)) air or g/h H2O equivalent | 22 | 3 |  |  | FALSE | Standard flow | 3.4028e+38 | -3.4028e+38 | NormMasFlw | 253 | FALSE | 113 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 1 | Controller speed factor (gain) | 30 | 8 | 3 |  | FALSE | Controller speed | 3.4028e+38 | 0 | Kspeed | 254 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Sensor code | 23 | 13 |  |  | TRUE | Sensor code | 65535 | 0 | SensorCode | 255 | FALSE | 113 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Sensor configuration code | 24 | 13 |  |  | TRUE | Sensor configuration code | 255 | 0 | SensorRevC | 256 | FALSE | 113 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Restriction code | 25 | 13 |  |  | TRUE | Restriction code | 65535 | 0 | RestrCode | 257 | FALSE | 113 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Restriction configuration code | 26 | 13 |  |  | TRUE | Restriction configurator code | 255 | 0 | RestrRevC | 258 | FALSE | 113 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Restriction NxP (proportional to air equivalent capacity of LFE) | 27 | 13 |  |  | TRUE | Restriction NxP | 2147483648 | 0 | RestrNxP | 259 | FALSE | 113 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | V,V | Seals information (1st byte = other, 2nd = plunger seal) | 28 | 13 |  |  | TRUE | Seals information | 255 | 0 | Seals | 260 | FALSE | 113 | TRUE | TRUE | 16 | c |  | TRUE
TRUE | TRUE | 0 | Valve code | 29 | 13 |  |  | TRUE | Valve code | 65535 | 0 | ValveCode | 261 | FALSE | 113 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Valve configuration code | 30 | 13 |  |  | TRUE | Valve configuration code | 255 | 0 | ValveRevC | 262 | FALSE | 113 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Instrument properties | 31 | 13 |  |  | TRUE | Instrument properties | 2147483648 | 0 | InstrProp | 263 | FALSE | 113 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 0 | Lookup table for frequency index | 10 | 3 |  |  | FALSE | Lookup table frequency index | 1 | 0 | LookFreqI | 264 | FALSE | 116 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Lookup table for frequency frequency | 11 | 3 |  |  | FALSE | Lookup table frequency frequency | 3.4028e+38 | -3.4028e+38 | LFFreq | 265 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Lookup table for frequency temperature | 12 | 3 |  |  | FALSE | Lookup table frequency temperature | 3.4028e+38 | -273.15 | LFTemp | 266 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Lookup table for frequency density | 13 | 3 |  |  | FALSE | Lookup table frequency density | 3.4028e+38 | -3.4028e+38 | LFDensity | 267 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Lookup table for frequency span adjust | 14 | 3 |  |  | FALSE | Lookup table frequency span adjust | 3.4028e+38 | -3.4028e+38 | LFSpanAdj | 268 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | Capacity (readout) unit index (extended unit table) | 15 | 3 | 19 |  | FALSE | Capacity unit index (ext) | 255 | 0 | CUnit | 269 | FALSE | 65 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | Actual density, measured by instrument | 15 | 3 |  |  | FALSE | Density actual | 3.4028e+38 | -3.4028e+38 | DensityAct | 270 | TRUE | 116 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Measured restriction | 18 | 13 |  |  | FALSE | Measured restriction | 3.4028e+38 | 0 | RestrMeas | 271 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Potmeter for sensor temperature compensation | 8 | 3 |  |  | FALSE | Temperature potmeter | 255 | 0 | TempPotm | 272 | FALSE | 116 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Gain for sensor temperature compensation | 9 | 3 |  |  | FALSE | Temperature potmeter gain | 255 | 0 | TempGain | 273 | FALSE | 116 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 1 | Counter controller overrun correction | 10 | 16 |  |  | FALSE | Counter controller overrun correction | 3.4028e+38 | 0 | CntrCConv | 274 | FALSE | 104 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 1 | Counter controller gain | 11 | 16 |  |  | FALSE | Counter controller gain | 3.4028e+38 | 0 | CntrCGain | 275 | FALSE | 104 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Sub fluid number | 1 | 3 |  |  | FALSE | Sub fluid number | 255 | 0 | FluidSub | 276 | FALSE | 65 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | Temperature compensation factor | 17 | 9 |  |  | FALSE | Temperature compensation factor | 3.4028e+38 | -3.4028e+38 | FreqTCor | 277 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | DSP register address | 29 | 13 |  |  | FALSE | DSP register address | 4294967295 | 0 | DSPRegI | 278 | FALSE | 116 | TRUE | FALSE |  | l | c | TRUE
TRUE | TRUE | 0 | DSP register long | 30 | 13 |  |  | FALSE | DSP register long | 4294967295 | -4294967295 | DSPRegLng | 279 | FALSE | 116 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 0 | DSP register floating point | 30 | 13 |  |  | FALSE | DSP register floating point | 3.4028e+38 | -3.4028e+38 | DSPRegFlt | 280 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | DSP register integer | 31 | 13 |  |  | FALSE | DSP register integer | 65535 | -65535 | DSPRegInt | 281 | FALSE | 116 | TRUE | TRUE |  | i |  | TRUE
FALSE | TRUE | 0 | Standard deviation | 0 | 19 |  |  | FALSE | Standard deviation | 3.4028e+38 | -3.4028e+38 | StdDev | 282 | TRUE | 121 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Measurement status | 1 | 19 |  |  | FALSE | Measurement status | 65535 | 0 | MeasStatus | 283 | FALSE | 121 | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | 0 | Measurement stop criteria | 2 | 19 |  |  | FALSE | Measurement stop criteria | 65535 | 0 | MStopCrit | 284 | FALSE | 121 | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | 0 | Measurement time out | 3 | 20 |  |  | FALSE | Measurement time out | 65535 | 0 | MTimeOut | 285 | FALSE | 121 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | Measurement maximum number of runs | 4 | 20 |  |  | FALSE | Maximum number of runs | 65535 | 0 | MMaxNrRuns | 286 | FALSE | 121 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | Measurement minimum standard deviation | 5 | 20 |  |  | FALSE | Minimum standard deviation | 3.4028e+38 | -3.4028e+38 | MMinStdDev | 287 | FALSE | 121 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | IO status for switches | 31 | 10 |  |  | FALSE | IO switch status | 4294967295 | 0 | IOSwitchSt | 288 | FALSE | 114 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 513 | Sensor bridge settings | 21 | 9 |  |  | TRUE | Sensor bridge settings | 65535 | 0 | SensBridge | 295 | FALSE | 65 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Sensor bridge current (0..0.02A) | 22 | 9 |  |  | TRUE | Sensor bridge current | 3.4028e+38 | 0 | SensCurren | 296 | FALSE | 65 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Sensor resistance (Ohm) | 23 | 9 |  |  | FALSE | Sensor resistance | 3.4028e+38 | 0 | SensResist | 297 | FALSE | 65 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Sensor bridge voltage (0..11.5V) | 24 | 9 |  |  | FALSE | Sensor bridge voltage | 3.4028e+38 | 0 | SensVolt | 298 | FALSE | 65 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE |  | Sensor name (based on sensor detection / sensor resistance) | 25 | 9 |  |  | FALSE | Sensor group name |  |  | SensName | 299 | FALSE | 65 | TRUE | FALSE | -2 | c |  | FALSE
TRUE | TRUE | 0 | Sensor temperature at calibration | 20 | 9 |  |  | FALSE | Sensor calibration temperature | 3.4028e+38 | -3.4028e+38 | SensCalTmp | 300 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | Valve safe state | 31 | 8 |  |  | FALSE | Valve safe state | 255 | 0 | ValveSafe | 301 | FALSE | 115 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Counter unit type | 12 | 16 |  |  | FALSE | Counter unit type index | 255 | 0 | CnUnitType | 302 | FALSE | 104 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE |  | Counter unit type name | 13 | 16 |  |  | FALSE | Counter unit type |  |  | CnUnTypNam | 303 | FALSE | 104 | TRUE | FALSE | 20 | c |  | TRUE
FALSE | TRUE | 0 | Counter unit index (extended counter unit table) | 14 | 16 |  |  | FALSE | Counter unit index (ext) | 255 | 0 | CnUnit | 304 | FALSE | 104 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 255 | Fieldbus1 select (main interface) | 8 | 23 |  |  | FALSE | Fieldbus1 selection | 255 | 0 | Bus1Select | 305 | FALSE | 125 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 1 | Fieldbus1 medium id (main interface) | 11 | 23 |  |  | FALSE | Fieldbus1 medium | 255 | 0 | Bus1Medium | 306 | FALSE | 125 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Fieldbus2 mode | 7 | 23 |  |  | FALSE | Fieldbus2 mode | 255 | 0 | Bus2Mode | 307 | FALSE | 124 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 2 | Fieldbus2 select | 8 | 23 |  |  | FALSE | Fieldbus2 selection | 255 | 0 | Bus2Select | 308 | FALSE | 124 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 2 | Fieldbus2 address | 10 | 23 |  |  | FALSE | Fieldbus2 address | 255 | 0 | Bus2Addr | 309 | FALSE | 124 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 12000000 | Fieldbus2 baudrate | 9 | 23 |  |  | FALSE | Fieldbus2 baudrate | 10000000000 | 0 | Bus2Baud | 310 | FALSE | 124 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 0 | Fieldbus2 medium id | 11 | 23 |  |  | FALSE | Fieldbus2 medium | 255 | 0 | Bus2Medium | 311 | FALSE | 124 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE |  | Fieldbus2 diagnose string | 20 | 23 |  |  | FALSE | Fieldbus2 diagnostics |  |  | Bus2Diagn | 312 | FALSE | 124 | TRUE | FALSE | -2 | c |  | FALSE
FALSE | TRUE | Propar | Fieldbus2 name | 21 | 23 |  |  | FALSE | Fieldbus2 name |  |  | Bus2Name | 313 | FALSE | 124 | TRUE | FALSE | -2 | c |  | FALSE
TRUE | TRUE | 0 | PIO channel selection | 0 | 10 |  |  | FALSE | PIO channel selection | 255 | 0 | PIOChSel | 314 | FALSE | 120 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | PIO parameter connected to i/o channel (=process number*256+parameter number) | 2 | 10 |  |  | FALSE | PIO parameter | 65535 | 0 | PIOParm | 315 | FALSE | 120 | TRUE | FALSE |  | i |  | TRUE
TRUE | TRUE | 0 | PIO input/output filter constant (0..1, 0=max, 1=min) | 6 | 10 |  |  | FALSE | PIO input/output filter | 3.4028e+38 | -3.4028e+38 | PIOFilter | 316 | FALSE | 120 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | PIO parameter capacity, 0% value | 7 | 10 |  |  | FALSE | PIO parameter capacity 0% | 3.4028e+38 | -3.4028e+38 | PIOPrmCap0 | 317 | FALSE | 120 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | PIO parameter capacity, 100% value | 3 | 10 |  |  | FALSE | PIO parameter capacity 100% | 3.4028e+38 | -3.4028e+38 | PIOPrmCap | 318 | FALSE | 120 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | PIO i/o channel configuration selection | 1 | 10 |  |  | FALSE | PIO configuration selection | 255 | 0 | PIOCfgSel | 319 | FALSE | 120 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | PIO analog input/output, zero adjust value | 4 | 10 |  |  | FALSE | PIO analog zero adjust | 3.4028e+38 | -3.4028e+38 | PIOAnZA | 320 | FALSE | 120 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | PIO analog input/output, span adjust factor | 5 | 10 |  |  | FALSE | PIO analog span adjust | 3.4028e+38 | -3.4028e+38 | PIOAnSA | 321 | FALSE | 120 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.) | 8 | 10 |  |  | FALSE | PIO hardware capacity max | 3.4028e+38 | -3.4028e+38 | PIOHwCapMx | 322 | FALSE | 120 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | PIO capacity set selection | 9 | 10 |  |  | FALSE | PIO capacity set selection | 255 | 0 | PIOCapSel | 323 | FALSE | 120 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.) | 10 | 10 |  |  | FALSE | PIO hardware capacity 0% | 3.4028e+38 | -3.4028e+38 | PIOHwCap0 | 324 | FALSE | 120 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.) | 11 | 10 |  |  | FALSE | PIO hardware capacity 100% | 3.4028e+38 | -3.4028e+38 | PIOHwCap | 325 | FALSE | 120 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Hardware platform identification number | 6 | 12 |  |  | FALSE | Hardware platform id | 65535 | 0 | HardwId | 326 | FALSE | 0 | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | 0 | Hardware platform sub identification number | 7 | 12 |  |  | FALSE | Hardware platform sub id | 255 | 0 | HardwSubId | 327 | FALSE | 0 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 0 | Temporary volatile fieldbus2 baudrate | 31 | 23 |  |  | FALSE | Temporary baudrate | 10000000000 | 0 | TempBaud | 328 | FALSE | 124 | TRUE | FALSE |  | l |  | TRUE
FALSE | TRUE | 0 | Monitor mode for setpoint | 23 | 2 |  |  | FALSE | Setpoint monitor mode | 255 | 0 | SetMonitor | 329 | FALSE | 115 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Special BHT parameter | 12 | 14 |  |  | FALSE | BHT12 | 10000000000 | 0 | BHT12 | 330 | FALSE | 118 | TRUE | FALSE |  | l |  | FALSE
TRUE | TRUE | 0 | Nominal sensor voltage (used for sensor temperature voltage compensation) | 26 | 9 |  |  | FALSE | Nominal sensor voltage | 3.4028e+38 | 0 | SenNomVolt | 331 | FALSE | 65 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Sensor voltage compensation factor (used for sensor temperature voltage comp.) | 16 | 9 |  |  | FALSE | Sensor voltage compensation factor | 3.4028e+38 | -3.4028e+38 | VoltTCor | 332 | FALSE | 116 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE |  | PCB serial number (unique number or string written by PCB manufacturer) | 31 | 4 |  |  | TRUE | PCB serial number |  |  | PCBSerial | 333 | FALSE | 119 | TRUE | TRUE | -2 | c |  | TRUE
TRUE | TRUE | 0 | Minimum measure time between sensors of a piston prover in 0.1 seconds | 24 | 20 |  |  | FALSE | Minimum measure time | 255 | 0 | MinMTime | 334 | FALSE | 115 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Fieldbus1 parity (main interface) | 12 | 23 |  |  | FALSE | Fieldbus1 parity | 255 | 0 | Bus1Parity | 335 | FALSE | 125 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | Fieldbus2 parity | 12 | 23 |  |  | FALSE | Fieldbus2 parity | 255 | 0 | Bus2Parity | 336 | FALSE | 124 | TRUE | TRUE |  | c |  | TRUE
FALSE | TRUE | 0 | Firmware identification number | 8 | 12 |  |  | FALSE | Firmware id | 255 | 0 | FirmwId | 337 | FALSE | 0 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 0 | Valve 100% current/voltage | 13 | 8 |  |  | FALSE | Valve 100% | 3.4028e+38 | 0 | Valve100 | 338 | FALSE | 114 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Minimum setpoint for controller | 14 | 8 |  |  | FALSE | Setpoint minimum | 1 | 0 | SetpMin | 339 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE |  | PCB serial number index | 30 | 4 | 0 | 0 | FALSE | PCB serial number index | 255 | 0 | PCBSerialI | 340 | FALSE | 119 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 00-00-00-00-00-00 | Fieldbus1 MAC address | 13 | 23 |  |  | FALSE | Fieldbus1 MAC address |  |  | Bus1MAC | 341 | FALSE | 125 | TRUE | FALSE | -2 | c |  | FALSE
TRUE | TRUE | 0 | Index of source fluidset for conversion | 0 | 3 |  |  | FALSE | Source fluidset index | 255 | 0 | FluidNrSrc | 342 | FALSE | 126 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Calculation method for capacity range | 1 | 3 |  |  | FALSE | Capacity range calculation method | 65535 | 0 | CapRanCalM | 343 | FALSE | 126 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Actual fluid property index | 2 | 3 |  |  | FALSE | Actual fluid property index | 255 | 0 | ActFlPIdx | 344 | FALSE | 126 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Actual fluid property calculation method | 3 | 3 |  |  | FALSE | Actual fluid property calculation method | 65535 | 0 | ActFlPCalM | 345 | FALSE | 126 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Mix fraction type | 4 | 3 |  |  | FALSE | Mix fraction type | 255 | 0 | MixFrcType | 346 | FALSE | 126 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Mix volume type temperature (�C) | 5 | 3 |  |  | FALSE | Mix volume temperature | 3.4028e+38 | -273.15 | MixVolTemp | 347 | FALSE | 126 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Mix volume type pressure (bar (a)) | 6 | 3 |  |  | FALSE | Mix volume pressure | 3.4028e+38 | 0 | MixVolPres | 348 | FALSE | 126 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | 0 | Mix component index | 7 | 3 |  |  | FALSE | Mix component index | 255 | 0 | MixCompIdx | 349 | FALSE | 126 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | Mix component fraction | 8 | 3 |  |  | FALSE | Mix component fraction | 1 | 0 | MixCompFrc | 350 | FALSE | 126 | TRUE | TRUE |  | f |  | TRUE
FALSE | TRUE | AIR | Mix component fluid name | 9 | 3 |  |  | FALSE | Mix component fluid name |  |  | MixCompNam | 351 | FALSE | 126 | TRUE | TRUE | 10 | c |  | TRUE
TRUE | TRUE | AIR | Fluid name for equivalent measure | 10 | 19 |  |  | FALSE | Equivalent measure fluid name |  |  | EqMFldName | 352 | FALSE | 126 | TRUE | FALSE | 10 | c |  | TRUE
TRUE | TRUE | 0 | Equivalent measure at actual conditions (kg/s) | 11 | 19 |  |  | FALSE | Equivalent measure | 3.4028e+38 | -3.4028e+38 | EqMeasure | 353 | FALSE | 126 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Fluid conversion conditions (temperature and pressure) selection | 12 | 3 |  |  | FALSE | Fluid conversion conditions selection | 65535 | 0 | FlConCondS | 354 | FALSE | 126 | TRUE | TRUE |  | i |  | TRUE
FALSE | TRUE | 0 | Sensor+restriction capacity minimum | 13 | 3 |  |  | FALSE | Sensor+restriction capacity minimum | 3.4028e+38 | -3.4028e+38 | SensCapMin | 355 | FALSE | 126 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Sensor+restriction capacity maximum | 14 | 3 |  |  | FALSE | Sensor+restriction capacity maximum | 3.4028e+38 | -3.4028e+38 | SensCapMax | 356 | FALSE | 126 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Valve capacity minimum | 15 | 3 |  |  | FALSE | Valve capacity minimum | 3.4028e+38 | -3.4028e+38 | ValvCapMin | 357 | FALSE | 126 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Valve capacity maximum | 16 | 3 |  |  | FALSE | Valve capacity maximum | 3.4028e+38 | -3.4028e+38 | ValvCapMax | 358 | FALSE | 126 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Maximum allowed dosing time (s) | 15 | 16 |  |  | FALSE | Maximum allowed dosing time | 3.4028e+38 | 0 | MaxDosTime | 359 | FALSE | 104 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Most recent dosing time (s) | 16 | 16 |  |  | FALSE | Most recent dosing time | 3.4028e+38 | 0 | DosingTime | 360 | FALSE | 104 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Controller hysteresis | 15 | 8 |  |  | FALSE | Controller hysteresis | 1 | 0 | CtrlHyst | 361 | FALSE | 114 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Pressure sensor input, zero adjust | 17 | 10 |  |  | FALSE | Pressure sensor input zero adjust | 3.4028e+38 | -3.4028e+38 | PressInZA | 362 | FALSE | 126 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 1 | Pressure sensor input, span adjust | 18 | 10 |  |  | FALSE | Pressure sensor input span adjust | 3.4028e+38 | -3.4028e+38 | PressInSA | 363 | FALSE | 126 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Special parameter index | 0 | 12 |  |  | FALSE | Special parameter index | 2147483648 | 0 | SpecialIdx | 364 | FALSE | 127 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 0 | Special parameter i long | 1 | 12 |  |  | FALSE | Special parameter i long | 2147483648 | -2147483648 | SpecLong | 365 | FALSE | 127 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 0 | Special parameter i float | 2 | 12 |  |  | FALSE | Special parameter i float | 3.4028e+38 | -3.4028e+38 | SpecFloat | 366 | FALSE | 127 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE |  | Special parameter i name | 3 | 12 |  |  | FALSE | Special parameter i name |  |  | SpecName | 367 | FALSE | 127 | TRUE | FALSE | -2 | c |  | TRUE
TRUE | TRUE | 0 | Special parameter l1 long | 4 | 12 |  |  | FALSE | Special parameter l1 long | 2147483648 | -2147483648 | SpecLong1 | 368 | FALSE | 127 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 0 | Special parameter l2 long | 5 | 12 |  |  | FALSE | Special parameter l2 long | 2147483648 | -2147483648 | SpecLong2 | 369 | FALSE | 127 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 0 | Special parameter l3 long | 6 | 12 |  |  | FALSE | Special parameter l3 long | 2147483648 | -2147483648 | SpecLong3 | 370 | FALSE | 127 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 0 | Special parameter l4 long | 7 | 12 |  |  | FALSE | Special parameter l4 long | 2147483648 | -2147483648 | SpecLong4 | 371 | FALSE | 127 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 0 | Special parameter l5 long | 8 | 12 |  |  | FALSE | Special parameter l5 long | 2147483648 | -2147483648 | SpecLong5 | 372 | FALSE | 127 | TRUE | FALSE |  | l |  | TRUE
TRUE | TRUE | 0 | Special parameter f1 float | 9 | 12 |  |  | FALSE | Special parameter f1 float | 3.4028e+38 | -3.4028e+38 | SpecFloat1 | 373 | FALSE | 127 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | Special parameter f2 float | 10 | 12 |  |  | FALSE | Special parameter f2 float | 3.4028e+38 | -3.4028e+38 | SpecFloat2 | 374 | FALSE | 127 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | Special parameter f3 float | 11 | 12 |  |  | FALSE | Special parameter f3 float | 3.4028e+38 | -3.4028e+38 | SpecFloat3 | 375 | FALSE | 127 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | Special parameter f4 float | 12 | 12 |  |  | FALSE | Special parameter f4 float | 3.4028e+38 | -3.4028e+38 | SpecFloat4 | 376 | FALSE | 127 | TRUE | FALSE |  | f |  | TRUE
TRUE | TRUE | 0 | Special parameter f5 float | 13 | 12 |  |  | FALSE | Special parameter f5 float | 3.4028e+38 | -3.4028e+38 | SpecFloat5 | 377 | FALSE | 127 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Fieldbus interface index | 7 | 23 |  |  | FALSE | Fieldbus interface index | 255 | 0 | BusIntIdx | 378 | FALSE | 125 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | Selftest status | 7 | 12 |  |  | FALSE | Selftest status | 2147483648 | 0 | SelfStat | 379 | FALSE | 119 | TRUE | FALSE |  | l |  | FALSE
TRUE | TRUE | 0 | Selftest block index | 8 | 12 |  |  | FALSE | Selftest block index | 255 | 0 | SelfBIdx | 380 | FALSE | 119 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | Selftest test index | 9 | 12 |  |  | FALSE | Selftest test index | 65535 | 0 | SelfTIdx | 381 | FALSE | 119 | TRUE | FALSE |  | i |  | TRUE
TRUE | TRUE | 0 | Selftest test description/name | 10 | 12 |  |  | FALSE | Selftest test description/name |  |  | SelfName | 382 | FALSE | 119 | TRUE | FALSE | -2 | c |  | FALSE
TRUE | TRUE | 0 | Selftest test result (pass/fail) | 11 | 12 |  |  | FALSE | Selftest test result | 255 | 0 | SelfRes | 383 | FALSE | 119 | TRUE | FALSE |  | c |  | FALSE
TRUE | TRUE | 0 | Selftest test lower limit | 12 | 12 |  |  | FALSE | Selftest test lower limit | 3.4028e+38 | -3.4028e+38 | SelfLower | 384 | FALSE | 119 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Selftest test upper limit | 13 | 12 |  |  | FALSE | Selftest test upper limit | 3.4028e+38 | -3.4028e+38 | SelfUpper | 385 | FALSE | 119 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Selftest test lowest value | 14 | 12 |  |  | FALSE | Selftest test lowest value | 3.4028e+38 | -3.4028e+38 | SelfLow | 386 | FALSE | 119 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Selftest test highest value | 15 | 12 |  |  | FALSE | Selftest test highest value | 3.4028e+38 | -3.4028e+38 | SelfHigh | 387 | FALSE | 119 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Selftest test average value | 16 | 12 |  |  | FALSE | Selftest test average value | 3.4028e+38 | -3.4028e+38 | SelfAvg | 388 | FALSE | 119 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Selftest test standard deviation | 17 | 12 |  |  | FALSE | Selftest test standard deviation | 3.4028e+38 | -3.4028e+38 | SelfStDv | 389 | FALSE | 119 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0.0.0.0 | IP address, when no address configured value is 0.0.0.0 (DHCP) | 14 | 23 |  |  | FALSE | Fieldbus1 IP address |  |  | Bus1IP | 390 | FALSE | 125 | TRUE | TRUE | -2 | c |  | TRUE
TRUE | TRUE | 0.0.0.0 | Subnet mask for the IP address | 15 | 23 |  |  | FALSE | Fieldbus1 subnet mask |  |  | Bus1SubMas | 391 | FALSE | 125 | TRUE | TRUE | -2 | c |  | TRUE
TRUE | TRUE | 0.0.0.0 | IP address of the gateway | 16 | 23 |  |  | FALSE | Fieldbus1 gateway address |  |  | Bus1GateW | 392 | FALSE | 125 | TRUE | TRUE | -2 | c |  | TRUE
FALSE | TRUE | 0 | Totalizer value | 17 | 16 |  |  | FALSE | Totalizer value | 3.4e+40 | -3.4e+40 | TotalzrVal | 393 | FALSE | 104 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | ln/min | Totalizer unit | 18 | 16 |  |  | FALSE | Totalizer unit |  |  | TotalzrUnt | 394 | FALSE | 104 | TRUE | TRUE | 4 | c |  | TRUE
FALSE | TRUE | 0 | Counter threshold | 20 | 16 |  |  | FALSE | Counter threshold | 3.4e+40 | -3.4e+40 | CntrThresh | 395 | FALSE | 104 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE |  | Contains fieldbus firmware version information | 22 | 23 |  |  | FALSE | Fieldbus1 version |  |  | Bus1Ver | 396 | FALSE | 125 | TRUE | FALSE | -2 | c |  | FALSE
TRUE | TRUE | 0 | Relative humidity (0..100%) | 20 | 9 |  |  | FALSE | Relative humidity | 100 | 0 | RelHum | 397 | FALSE | 65 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Dosing type | 1 | 16 |  |  | TRUE | Dosing type | 255 | 0 | DosType | 398 | FALSE | 112 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Dosing controller type | 2 | 16 | 0 | 0 | TRUE | Dosing controller type | 255 | 0 | DosCtrType | 399 | FALSE | 112 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Dosing rejection type | 3 | 16 | 0 | 0 | FALSE | Dosing rejection mode | 255 | 0 | DosActType | 400 | FALSE | 112 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Dosing mode | 4 | 16 | 0 | 0 | FALSE | Dosing mode | 255 | 0 | DosStartTp | 401 | FALSE | 112 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Dosing start delay time | 5 | 16 | 0 | 0 | FALSE | Dosing start delay time | 3.4e+40 | 0 | DosDelayT | 402 | FALSE | 112 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Dosing time | 6 | 16 | 0 | 0 | FALSE | Dosing time | 3.4e+40 | 0 | DosTime | 403 | FALSE | 112 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Dosing repetition interval | 7 | 16 | 0 | 0 | FALSE | Dosing repetition interval | 3.4e+40 | 0 | DosRepInt | 404 | FALSE | 112 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Dosing size | 8 | 16 | 0 | 0 | FALSE | Dosing size | 3.4e+40 | 0 | DosSize | 405 | FALSE | 112 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Dosing accuracy [%] | 9 | 16 | 0 | 0 | FALSE | Dosing accuracy [%] | 3.4e+40 | 0 | DosAcc | 406 | FALSE | 112 | TRUE | FALSE |  | f |  | TRUE
FALSE | TRUE | 0 | Actual dosing size | 10 | 16 | 0 | 0 | FALSE | Actual dosing size | 3.4e+40 | 0 | ActDosSze | 407 | TRUE | 112 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Actual dosing dosing time [s] | 11 | 16 | 0 | 0 | FALSE | Actual dosing time [s] | 3.4e+40 | 0 | ActDosTime | 408 | TRUE | 112 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Actual dosing accuracy [%] | 12 | 16 | 0 | 0 | FALSE | Actual dosing accuracy [%] | 3.4e+40 | -3.4e+40 | ActDosAcc | 409 | TRUE | 112 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Dosing unit | 0 | 16 | 0 | 0 | FALSE | Dosing unit |  |  | DosUnt | 410 | FALSE | 112 | TRUE | FALSE | 4 | c |  | TRUE
FALSE | TRUE | 0 | Diagnostic newest event index | 14 | 24 | 0 | 0 | FALSE | Diagnostic newest event index | 65535 | 0 | DiaNEvtIdx | 411 | FALSE | 118 | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | 0 | Diagnostic event index | 15 | 24 | 0 | 0 | FALSE | Diagnostic event index | 65535 | 0 | DiaEvtIdx | 412 | FALSE | 118 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | Diagnostic event code | 16 | 24 | 0 | 0 | FALSE | Diagnostic event code | 65535 | 0 | DiaEvtCode | 413 | FALSE | 118 | TRUE | FALSE |  | i |  | FALSE
FALSE | TRUE | 0 | Diagnostic event description | 20 | 24 | 0 | 0 | FALSE | Diagnostic event description |  |  | DiaEvtDesc | 414 | FALSE | 118 | TRUE | FALSE | -2 | c |  | FALSE
FALSE | TRUE | 0 | Diagnostic event active | 17 | 24 | 0 | 0 | FALSE | Diagnostic event active | 255 | 0 | DiaEvtAct | 415 | FALSE | 118 | TRUE | FALSE |  | c |  | FALSE
FALSE | TRUE | 0 | Diagnostic event namur status | 18 | 24 | 0 | 0 | FALSE | Diagnostic event namur status | 255 | 0 | DiaEvtNSt | 416 | FALSE | 118 | TRUE | FALSE |  | c |  | FALSE
FALSE | TRUE | 0 | Diagnostic event timestamp | 21 | 24 | 0 | 0 | FALSE | Diagnostic event timestamp | 4294967295 | 0 | DiaEvtTime | 417 | FALSE | 118 | TRUE | FALSE |  | l |  | FALSE
FALSE | TRUE | 0 | Instrument namur status | 0 | 24 | 0 | 0 | FALSE | Instrument namur status | 255 | 0 | InstrNSt | 418 | FALSE | 118 | TRUE | FALSE |  | c |  | FALSE
FALSE | TRUE | 0 | Instrument timestamp | 13 | 24 | 0 | 0 | FALSE | Instrument timestamp | 4294967295 | 0 | InstrTime | 419 | FALSE | 118 | TRUE | TRUE |  | l |  | TRUE
FALSE | TRUE | 0 | Operational history parameter index | 26 | 24 | 0 | 0 | FALSE | Operational history parameter index | 65535 | 0 | OpHisIdx | 420 | FALSE | 118 | TRUE | FALSE |  | i |  | TRUE
FALSE | TRUE | 0 | Operational history parameter name | 27 | 24 | 0 | 0 | FALSE | Operational history parameter name |  |  | OpHisName | 421 | FALSE | 118 | TRUE | FALSE | -2 | c |  | FALSE
FALSE | TRUE | 0 | Operational history parameter minimum value | 28 | 24 | 0 | 0 | FALSE | Operational history parameter minimum value | 3.4e+40 | 0 | OpHisMin | 422 | FALSE | 118 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Operational history parameter maximum value | 29 | 24 | 0 | 0 | FALSE | Operational history parameter maximum value | 3.4e+40 | 0 | OpHisMax | 423 | FALSE | 118 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Operational history parameter average | 30 | 24 | 0 | 0 | FALSE | Operational history parameter average | 3.4e+40 | 0 | OpHisAvg | 424 | FALSE | 118 | TRUE | FALSE |  | f |  | FALSE
FALSE | TRUE | 0 | Operational history parameter standard deviation | 31 | 24 | 0 | 0 | FALSE | Operational history parameter standard deviation | 3.4e+40 | 0 | OpHisStdDv | 425 | FALSE | 118 | TRUE | FALSE |  | f |  | FALSE
TRUE | TRUE | 0 | Fieldbus stop bits | 23 | 23 | 0 | 0 | FALSE | Fieldbus stop bits | 255 | 0 | Bus1StopB | 426 | FALSE | 125 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Fieldbus connection mode | 24 | 23 | 0 | 0 | FALSE | Fieldbus connection mode | 255 | 0 | Bus1ConMod | 427 | FALSE | 125 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Fieldbus passkey | 25 | 23 | 0 | 0 | FALSE | Fieldbus passkey | 4294967295 | 0 | Bus1PssKey | 428 | FALSE | 125 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE |  | Feature unlock code | 27 | 12 | 0 | 0 | FALSE | Feature unlock code |  |  | FeatUnlckC | 429 | FALSE | 119 | TRUE | FALSE | -2 | c |  | TRUE
TRUE | TRUE | 0 | Feature selection | 28 | 12 | 0 | 0 | FALSE | Feature selection | 255 | 0 | FeatSel | 430 | FALSE | 119 | TRUE | FALSE |  | c |  | TRUE
TRUE | TRUE | 0 | Feature state | 29 | 12 | 0 | 0 | FALSE | Feature state | 255 | 0 | FeatState | 431 | FALSE | 119 | TRUE | FALSE |  | c |  | FALSE
FALSE | TRUE | 0 | Control function | 10 | 8 | 0 | 0 | FALSE | Control function | 255 | 0 | CtrlFunc | 432 | FALSE | 115 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE | 0 | Actuator index | 0 | 8 | 0 | 0 | FALSE | Actuator index | 255 | 0 | ActIdx | 433 | FALSE | 114 | TRUE | FALSE |  | c |  | TRUE
FALSE | TRUE |  | Dosing status | 13 | 24 | 0 | 0 | FALSE | Dosing status | 65535 | 0 | DosStatus | 434 | FALSE | 112 | TRUE | FALSE |  | i |  | FALSE
TRUE | TRUE | 0 | Body code | 9 | 13 |  |  | TRUE | Body code | 65535 | 0 | BodyCode | 435 | FALSE | 114 | TRUE | TRUE |  | i |  | TRUE
TRUE | TRUE | 0 | Body configuration code | 10 | 13 |  |  | TRUE | Body configuration code | 255 | 0 | BodyRevCod | 436 | FALSE | 114 | TRUE | TRUE |  | c |  | TRUE
TRUE | TRUE | 0 | Batch sequence number | 14 | 16 |  |  | TRUE | Batch sequence number | 4294967295 | 0 | DosSeqNr | 437 | FALSE | 112 | TRUE | TRUE |  | l |  | TRUE
TRUE | TRUE | 0 | Lookup table for linearisation delta p | 27 | 3 |  |  | FALSE | Lookup table delta p | 3.4028e+38 | -3.4028e+38 | LookDp | 438 | FALSE | 65 | TRUE | TRUE |  | f |  | TRUE
TRUE | TRUE | 0 | Lookup table for linearisation pressure | 28 | 3 |  |  | FALSE | Lookup table pressure | 3.4028e+38 | -3.4028e+38 | LookP | 439 | FALSE | 65 | TRUE | TRUE |  | f |  | TRUE